DEPTH for Windows (3.0)

To calculate DEPTH RESOLUTION and to SIMULATE SPECTRA

by E. Szilágyi

Using this program you can calculate:

(1) the depth resolution for Elastic Recoil Detection Analysis (ERD), Backscattering Spectrometry (BS) and Forward Scattering (FS) for a given geometry; In this version the calculation of Nuclear Reaction Analysis (NRA) is not yet included.

(2) the optimal depth resolution;

(3) the energy loss and the Bohr straggling of the incident ion (at the incident energy) going through the absorber foil;

(4) the Tschalar effect without and with a given inhomogeneity;

(5) sub-spectra (of a given isotope) accounting for all energy-spread effects of RBS and ERDA;

(6) full spectra accounting for all energy-spread effects of RBS and ERDA;

(7) resonance method.

In calculations of the energy-spread values for the energy loss methods the following contributions (all in FWHM sense) are taken into account based on ref. [5]:

– The energy and angular spread of the beam. These contributions are connected to the energy stability of the accelerator and to the beam handling system, respectively.

– The geometrical spread caused by the finite beam spot size and detector solid angle.

– The straggling of incident ions and emitted particles in the target. Here both the nuclear and electronic straggling contributions are treated. The latter can be calculated both with and without Chu's correction. The charge state fluctuations for heavy ions are treated using Yang’s formulae. We also take into account the effect of change in stopping power for large energy losses. We assume the straggling contributions to be Gaussian-shaped and the particles' paths in and out produce independent effects.

– The multiple scattering of incident ions and emitted particles in the target, where only projected distributions play a role. The multiple scattering causes energy spread in two ways:

(i) the energy of the ions changes because of the different path lengths, and

(ii) the energy of the emitted particle depends on the reaction angle.

The combination of the last two effects is treated as statistically dependent, while the ways in and out are independent. The shape of the resulting energy spread was approximated with Pearson VII distribution (The Gaussian and the Breit - Wigner distributions are special cases derived from Pearson VII.) Applying a special procedure, we are able to treat multi-elemental samples and large values of energy losses as well.

– The energy resolution of the detecting method, it is taken into account with Gaussian shape. The code is able to handle the surface barrier detectors and the magnetic spectrographs [9].

– The absorber foil (if there is any). The width of the detected energy spread measured with an absorber foil is obtained in two steps. First we calculate the width and shape of the energy distribution just at the exit from the sample, then using the transfer function of the foil we determine the correct detected energy spectrum and its FWHM using numerical integration.

Numerical convolution of the above mentioned independent contributions would be carried out to obtain the total energy distribution before the absorber foil. This would be, however, a relatively slow procedure. To avoid this, an approximate fast summation method was worked out for Pearson VII distribution.

The following effects are not included in our calculations: plural scattering, extremely small energy losses, non-Gaussian straggling for small depths and sample surface roughness [2].

The Wdepth.zip file contains the following directories and files:

| Directory | file name | Remarks |
| --- | --- | --- |
| c:\wdepth\ | Wdepth.exe input.asc Wdepth.doc | the code sample input file (Word for Windows 6.0), this file |
| c:\Wdepth\lev\ | chupar.dat fmt.dat foils.dat reactions.dat stopping.dat taufi.dat taufij.dat *.lev f??.dat | Fitted parameter values for Chu’s correction Fitted parameters for convolution of Pearson VII function Identify the different, already calculated absorber foils Q value of the reactions for the ground state Stopping power data for multiple scattering calculation for multiple scattering calculation level structures for different nuclei: e.g. O16.lev energy straggling of the absorber foils |
| c:\Wdepth\Cross_sections\ | *.dat | Gurbich's cross section data for RBS. |
| c:\WDepth\ZBL2003\ | *.dat | Stopping tabels for selected ions. |
| c:\Wdepth\test |  | Examples for selected problems (including parameter files, and results, etc. |
| Parameter files, layer descriptions, results can be saved to any directory. |  |  |

How to use the program:

To start the program you can create a shortcut to your desktop or simply run wdepth.exe.

First the parameters have to define; either by loading a parameter file (file menu / load a new parameter file) or by setting the individual parameters by selecting "parameters" in the menu line.

After setting all the necessary parameters for a given calculations, the parameter files can be saved. The parameter files do not contain the description of samples and absorbents.

The code can be run from a DOS-window using input file, too. In this case a predefined parameter files are necessary. The input file is a simple ASCII file, in which you have to give in separate lines where the parameter, target and absorber files are, which calculation will perform, and the file name for the results, e.g.,:

c:\wdepth\test\HeERD_AG\Heerd.par

c:\wdepth\test\HeERD_AG\AgERD.lay

c:\wdepth\test\HeERD_AG\mylar.lay

depth and energy resolution

c:\wdepth\test\HeERD_AG\HeerdAG.dat

Setting the parameters:

The parameter set divided to several tabs as beam, sample, detection, options, optimisations, stopping powers, cross section and simulations. Each parameter begins with a text description, which defines the parameter and an editable field.

For the depth resolution calculation (or simulating a sub-spectrum) the target element and its mass number specified in the reaction part are used, its abundance is used only for the spectrum calculation. For the elastic scattering the Qm value and the excitation level should be equal to 0.

Samples and absorbers:

Sample description always starts with the thickness of the layer (expressed in A or the symbol of the substrate, S), the next line contains the density of the layer in at/cm3, then in separate line(s) the chemical symbol of the elements, and the relative concentrations of the atoms e.g., :

remarks

a=50	thickness in Angstrom

n=1e23	density of the layer in at/cm3

c,1	chemical symbol of a given element (most abundant isotope), relative concentrations of the atoms

h,1

a=13

n=1e23

si#,1	#: to calculate all the natural isotopes of the elements

O18,2	calculate only a specified isotope

A=37000

N=4.978e22

Si#,1

s=	substrate

N=4.978e22

Si#,1

The absorber file is similar, but without substrate. Both file can be saved as *.lay file.

Detector position:

IBM – if the detector is mounted in the plane determined by the beam and the sample surface normal

CORNELL – if the detector is mounted in the plane determined by the beam and the sample tilt axis.

Magnetic spectrograph and surface burrier detectors are considered. The energy resolution of the detector, dEdet is fitted using the following function:

dEdet = a0 + a1 Ea2 + a3 Ea4

a0, a1, a2, a3 and a4 parameters can be given by clicking to parameters, detection and energy resolution.

The energy calibration of the detector is used only for simulation. Linear energy calibration (keV/ch and E0) can be given by clicking to parameters, detection and energy calibration.

In the depth resolution calculation rough error estimation is given. If the theory described below is not valid anymore the program gives a total depth resolution equal to 10000 nm. The program gives the maximal probing depth, when the mean detected energy of the recoils will be smaller than the threshold energy.

Cross sections

Rutherford cross-section can be used with or without screening. Special fits can be used for ERD (He, p) and (He, d) reactions. For the cross section of p(4He,p)4He reaction the Benenson fit's is described in ref. [6], Quillet's cross section and Szilágyi's cross section are also included. Cross sections in ASCII format can be included. In the cross section files the energy and the cross section values have to be listed in separate lines (keV and mb/sr (in lab system) units). When the cross section files are imported, the scattering angles and the reaction details have to be defined.

Stopping powers:

Default stopping power is calculated by SRIM95. For selected ions SRIM03 stopping powers are tabulated. In lack of data, SRIM95 is used.

Calculations

Before starting your calculation check all the parameters carefully. Then choose the requested calculations from Calculations menu.

- energy loss:	calculates the energy loss of incident ions in the absorbent foil.

- Tschalar effect:	calculates the Tschalar-effect for the emitted ions in the absorbent foil.

- Depth and energy resolution:	calculates depth and energy resolution for the given geometry

- optimisation:	determine the best depth resolution by changing the geometry

- spectrum simulation:	calculates sub-spectrum for the given target isotope

- full spectrum: 	calculates the full spectrum

- exitation curves for resonance:	calculates excitation curves of the resonance method (yield vs. energy)

When the program finishes, the results can be saved clicking to "save" button of the "Output" panel or can be drawn clicking to "View results" button.

In case of depth resolution, the program may give one of the following messages:

– E < dE/10: the detected mean energy is smaller than the tenth of the detected energy spread

– absorber: the FWHM of the detected energy spread cannot be evaluated because it is too close to zero.

– detector resolution reaches 0.

– ctg : if somewhere: abs(/2 ctg )>1 meaning that the approximation used is not valid any more

– error>500%: meaning that the approximation used is not valid any more.

– stopping power = 0

– sample thickness!!! : The ions reach the end of the sample.

– Max. depth = ..... : The ions reach the maximal depth value.

– e < thr : the energy of the ions reaches the threshold energy

– the number of sublayers reaches the limit (150)

Output format

(1) Energy loss of the incident ion (at the incident energy) going through the absorber foil.

This option is useful to calculate the thickness of the absorber foil for ERDA or calibration of NRA spectra, (the result can be saved) or calculate the stopping power as a function of energy or depth.

/*  depth   Energy     Stopping power

/* [micron]  [keV]      [keV/micron]

0.0000   1620.0000   228.3863

0.0500   1608.5807   229.1652

0.1000   1597.1224   229.9507

0.1500   1585.6249   230.7430

0.2000   1574.0877   231.5420

0.2500   1562.5106   232.3477

The first column here is the depth in micron, the second one is the energy in keV and third one is the stopping power in keV/micron.

(4) Tschalar effect with and without a given inhomogeneity;

The energy straggling of the absorber foil can be calculated taking into account the effect of change in stopping power for large energy losses. This option can be used to estimate inhomogeneity of the absorber foil (see ref. 5). (Similar calculation is made automatically for the absorbers.)

The straggling values are calculated in two energy ranges with given energy steps.

Output format

e.g.: for 24 micron Mylar foil: E1 = 6000 keV, step1 = 250 keV and E2= 2000 , step2=50 we get:

/*  E             rel.E    de (without) de (with)

/*                              inhomogeneity

/* [keV]       -            [keV]         [keV]

6000.000      0.037     39.559     40.657

5750.000      0.040     39.620     40.797

5500.000      0.043     39.689     40.955

5250.000      0.047     39.767     41.133

5000.000      0.052     39.856     41.335

4750.000      0.057     39.958     41.567

4500.000      0.063     40.077     41.835

4250.000      0.069     40.216     42.146

4000.000      0.077     40.380     42.512

3750.000      0.087     40.575     42.946

3500.000      0.099     40.812     43.469

3250.000      0.113     41.103     44.108

3000.000      0.131     41.469     44.905

2750.000      0.154     41.939     45.923

2500.000      0.183     42.567     47.266

2000.000      0.279     44.785     51.873

1950.000      0.293     45.143     52.593

1900.000      0.308     45.543     53.392

1850.000      0.324     45.996     54.287

1800.000      0.342     46.512     55.295

1750.000      0.362     47.106     56.443

1700.000      0.383     47.800     57.765

1650.000      0.407     48.622     59.306

1600.000      0.434     49.614     61.134

1550.000      0.463     50.836     63.343

1500.000      0.497     52.386     66.084

1450.000      0.535     54.418     69.597

1400.000      0.580     57.210     74.302

1350.000      0.633     61.288     80.994

1300.000      0.697     67.777     91.366

The first column here is the incident energy in keV, the second one is the relative energy loss – the ion energy after the foil / incident energy and the last two column contain the straggling in the foil without and with inhomogeneity in keV.

(3) Depth and energy resolution

After the depth resolution calculation the results can be saved. Depth will create two files (*_1.dat and *_2.dat). The first one:

REFLECTION GEOMETRY

theta = 165.00         alpha = 80.00

depth| depth & energy  |      before the    |                main energy spread contributions          |error| mean

|   resolution    |       reaction       |                        after the sample                             |from | det.

x     |    dx    det      |   E      dE    dfi   | dopll.  beame   beama    geo     ss        msi      mso      tot    |appr.| energy

[nm]  |   [nm]  [keV]  |     [keV]    [deg]  |                          [keV]                                            |  %   | [keV]

0.0    28.52    17.4  3100.0  1.2  0.0     0.1     0.3     0.1     6.9     0.0     0.0     0.0     6.9   -0.9   792.8

50.0   29.98    18.3  3085.5  5.7  0.2     0.1     0.3     0.1     6.6     5.5     0.4     1.6     8.8   -0.8   762.3

.

.

.

1000.0   54.74  28.2  2801.9  26.6 1.2     0.1     0.2     0.1     3.4    18.2     2.1    11.8    23.1    0.2   180.0

1050.0   57.84  28.1  2786.5  27.3 1.2     0.1     0.2     0.1     3.4    17.4     2.0    12.4    22.5    0.2   155.7

the max. probing depth = 1050 nm

The calculation was finished: e < thr  column numbers:

1     |    2      3       |  4         5       6   |     7        8       9        10      11      12      13      14  |     15     |  16

The meaning of the columns are, as follows:

1)	the depth in nm units,

2)	the depth resolution in nm units,

3)	the energy resolution of the detected ions in keV,

4)	the mean energy of the incident ions before the reaction [keV],

5)	the energy spread of the incident ions before the reaction [keV],

6)	the angular spread of the incident ions before the reaction [deg],

Then the calculated energy spread contributions are listed, just after the ions leave the sample (before the detector and absorber foil) taking into account all the necessary corrections.

7)	energy spread of the beam [keV],

8)	energy spread caused by the angular spread of the beam [keV],

9)	geometrical energy spread [keV],

10)	straggling in the sample (inward and outward) [keV],

11)	energy spread caused by the multiple scattering effects (lateral and angular), inward [keV],

12)	energy spread caused by the multiple scattering effects (lateral and angular), outward [keV],

13)	energy spread caused by the doppler effect [keV],

14)	the “sum” of these contributions [keV],

Then:

15)	error coming from the main approximation [%],

16)	the mean energy after the particles leave the sample, [keV].

In the second file details on calculation of the multiple scattering is given:

alpha_in =  10.00     theta  = 20.00

beta_out_1 = 10.00  beta_out_2 = 90.00

Z1 = 2    M1 = 4    Z2 = 1    M2 = 1

depth |  straggling  | ms in after the sample |   ms out after the sample        |S eff    |m

|  in   |   out   | tau  | fi   |  nu  |  k  |  tau  |  fi       | nu   |  k    |  k      |           |res

[nm]  |     [keV]      |  -    |[deg]|  -   | - Á |   -     | [deg] |   -    | - Á   |  -º     | keV/nm]| -

0.0    0.0    0.0    0.0    0.0    2.0    0.0    0.0    0.0    2.0    0.0    0.0      1.3     100.0

5.0    3.1    1.5    1.3    0.2    1.2    0.0    1.4    0.1    1.2    0.0    0.0      1.3      17.1

10.0    4.4    2.2    2.6    0.3    1.3    0.0    2.8    0.2    1.3    0.0    0.0      1.3      13.7

15.0    5.4    2.7    3.8    0.4    1.3    0.0    4.2    0.3    1.4    0.0    0.0      1.3      12.7

1   |    2      3    |  4        5       6       7 |     8       9     10    11      12      | 13        14  |

1)	the depth in nm units,

2)	straggling inward [keV],

3)	straggling outward [keV],

the multiple scattering data for the inward path: for details see ref. 5

4)	tau - the reduced path length,

5)	fi - FWHM of the angular spread,

6)	nu(tau) - see: eq. 45 in ref. 1. ,

7)	k-perpendicular (the ratio of the lateral and angular spreads),

the multiple scattering data for the outward path

8)	tau - the reduced path length,

9)	fi,

10)	nu,

11)	k-perpendicular,

12)	k-parallel,

13)	effective stopping power [keV/nm],

14) 	m parameter of the energy spread distribution using Pearson VII approximation. (If m>35 the distribution can be considered as Gaussian distribution.)

(4) Optimal depth resolution

The optimal depth resolution as a function of the depth can be calculated using different geometries. If the detector position is IBM or CORNELL the program changes automatically the angles with a given step between minimal and maximal values of the scattering and incoming angles. If the detector angle is fixed you can give 4 different depths, for which the program saves the depth resolution values as a function of the incoming angle. In this case you have to give depth values divisible by the step given.

Output format

a) optimalization: (example for ERD: Theta min = 10 Theta max = 20 ; alpha min =5; step 5 degrees)

depth   depthres   alfa    theta

[nm]      [nm]     [deg]   [deg]

0.000    4.375    1.000    5.000

5.000    5.501    1.000    5.000

10.000    8.499    2.000    5.000

15.000    9.513    2.000    5.000

20.000   10.664    2.000    5.000

25.000   11.850    3.000    5.000

The first column here is the depth in nm, the second one is the depth resolution in nm, the next two column the geometry (alpha and theta, both in degrees) and the last one the error value.

b) at fixed scattering angle: (example for ERD: Theta min = Theta max = 20 ; alpha min =0.2; step 0.2 degrees)

alfa     theta          depth res. [nm] at

[deg]    [deg]          x1          x2            x3        x4

0.200   20.000   11.787    0.000    0.000    0.000

0.400   20.000   11.508    0.000    0.000    0.000

0.600   20.000   11.462    0.000    0.000    0.000

0.800   20.000   11.615    0.000    0.000    0.000

1.000   20.000   11.520    0.000    0.000    0.000

1.200   20.000   11.997    0.000    0.000    0.000

1.400   20.000   12.342    0.000    0.000    0.000

1.600   20.000   12.998   33.504    0.000    0.000

1.800   20.000   13.932   28.372    0.000    0.000

2.000   20.000   14.207   24.500    0.000    0.000

2.200   20.000   15.568   22.352    0.000    0.000

2.400   20.000   16.069   20.968    0.000    0.000

2.600   20.000   16.613   19.977   46.111    0.000

2.800   20.000   17.191   20.228   40.894    0.000

3.000   20.000   17.797   20.138   36.522    0.000

3.200   20.000   18.426   19.895   33.993    0.000

3.400   20.000   19.073   20.877   31.392    0.000

3.600   20.000   21.340   21.977   29.471    0.000

3.800   20.000   22.091   22.397   29.713   47.545

The first two columns contain the geometry (alpha and theta), the next 4 ones the depth resolution values at the given depths.

(5-6) Spectrum and full spectrum calculate spectra accounting for all energy-spread effects of RBS and ERDA. The results can be saved in two different formats. Depth files (*.txt), spectrum files (*.dat).

Depth files (here the experiment spectrum and its simulation can be saved into a single ASCII file:

channels     energy_[keV]     measured       simulated

1           18250.3                       0             33.1

2          18250.5                        0             33.2

3          18250.8                        0             33.3

4          18251.0                        0             33.4

5          18251.3                        0             33.6

The first column here the channels, detected energy, the experimental and the simulated yield.

spectrum files (only the results of the simulation are saved):

1            337.0

2            335.3

3            334.8

4            332.1

5            324.6

6            314.0

7            304.7

The first column here the channels, while the second one is the simulated yield.

(7) excitation curves for resonance: In the resonance method the beam energy increasing with the given step. In this example the 4260 keV resonance of the carbon was used. Below the resonance energy the depth resolution is not calculated.

REFLECTION GEOMETRY

theta = 165.00         alpha = 80.00      FWHM [keV] = 20.00

|  depth |  E_in  |               main energy spread contributions              | depth  |counts

|            |           |                  at the resonance energy                          |  res.  |

|    x      |            |   dopl.   beame     beama     ss    msi     tot    tot+FWHM|        |

|  [nm]   | [keV]  |                            [keV]                                              |  [nm]  |

0.0   4250.0      0.0      0.0      0.0      0.0      0.0      0.0      0.0         0.0          396799.5

0.0   4260.0      0.6      1.0      0.0      0.0      0.0      1.2     19.9        82.8        438072.2

41.7   4270.0      0.6      1.0      0.0      4.4      0.0      4.7     21.1        87.9        480286.9

1     |    2        |     3          4        5          6         7         8       9        |     10       |      11

The meaning of the columns are, as follows:

1)	the depth in nm units, (where the resonance criterium is fullfil).

2)	the energy of the incident ions in keV,

Then the main energy spread contributions are listed, at the resonance energy, all in keV units.

3)	energy spread due to Doppler contribution [keV],

4)	energy spread of the beam [keV],

5)	energy spread caused by the angular spread of the beam [keV],

6)	straggling in the sample (inward) [keV],

7)	energy spread caused by the multiple scattering effects (lateral and angular), inward [keV],

8)	all the energy spread contributions

9)	all the energy spread contributions + the FWHM of the resonance

10)	depth resolution of the resonance method

11)	peak integral of the resonance

References:

There are several publications related to this code. The base of the depth resolution calculation is described in ref 5.

# F. Pászti, E. Szilágyi, E. Kótai:

Optimization of the depth resolution in elastic recoil detection

Nucl. Instr. Meth. B 54 (1991) 507 - 512.

# N.P. Barradas, J.C. Soares, M.F. da Silva, E. Szilágyi, F. Pászti:

Study of multilayer interface roughness using RBS with improved depth resolution

Nucl. Instr. Meth. B 94 (1994) 266 - 270.

# E. Szilágyi, F. Pászti:

Theoretical calculation of the depth resolution of IBA methods

Nucl. Instr. Meth. B 85 (1994) 616 - 620.

# E. Szilágyi, F. Pászti, V. Quillet, F. Abel:

Optimization of the depth resolution in ERDA of H using 12C ions

Nucl. Instr. Meth. B 85 (1994) 63 - 67.

# E. Szilágyi, F. Pászti, G. Amsel:

Theoretical approximations for depth resolution calculations in IBA methods

Nucl. Instr. Meth. B 100 (1995) 103 - 121.

# E. Szilágyi, L.S. Wielunski, F. Pászti:

Theoretical Approximation of Energy Distribution of Elastically Recoiled Hydrogen Atoms

Nucl. Instr. Meth. B 136-138 (1998) 701 - 706.

# L.S. Wielunski, E. Szilágyi, G.L. Harding:

Multiple Scattering Effects in Elastic Recoil Detection Depth Resolution

Nucl. Instr. Meth. B 136-138 (1998) 713 - 718.

# E. Szilágyi:

Energy spread in ion beam analysis

Nucl. Instr. Meth. B 161-163 (2000) 37 - 47.

# E. Szilágyi:

On the limitations introduced by energy spread in elastic recoil detection analysis

Nucl. Instr. Meth. B 183 (2001) 25 - 33.

# E. Szilágyi, E. Kótai, Z. Zolnai, G. Battistig, T. Lohner, J. Gyulai:

Ion implantation induced damage in silicon carbide studied by non-Rutherford elastic backscattering

2000 International Conference on Ion Implantation Technology (IIT'2000) Proceedings, Alpbach, Austria, 17-22 September 2000, (Eds.: H. Ryssel, L. Frey, J. Gyulai, H. Glawishing), IEEE 00EX432, ISBN 0-7803-6462-7, 2000, pp.131-134.

# E. Szilágyi, N.Q. Khánh, Z.E. Horváth, T. Lohner, G. Battistig, Z. Zolnai, E. Kótai, J. Gyulai:

Ion Bombardment Induced Damage in Silicon Carbide Studied by Ion Beam Analytical Methods

Mater. Sci. Forum 353-356 (2001) 271 - 274.

# L.S. Wielunski, D. Grambole, U. Kreissig, R. Grötzschel, G. Harding, E. Szilágyi:

Hydrogen depth resolution in multilayer metal structures, comparison of elastic recoil detection and resonant nuclear reaction method

Nucl. Instr. Meth. B 190 (2002) 693 - 698.

# E. Rauhala, N.P. Barradas, S. Fazinic, M. Mayer, E. Szilágyi, M. Thompson:

Status of ion beam data analysis and simulation software

Nucl. Instr. Meth. B 244 (2006) 436 - 456.

# N.P. Barradas, K. Arstila, G. Battistig, M. Bianconi, N. Dytlewski, C. Jeynes, E. Kótai, G. Lulli, M. Mayer, E. Rauhala, E. Szilágyi, M. Thompson:

International Atomic Energy Agency intercomparison of Ion Beam Analysis software

Nucl. Instr. and Meth. in Phys. Res. B (2007),

If you have any problem or question using this program please feel free to contact me at:

Edit Szilágyi

MTA KFKI Research Institute for Particle and Nuclear Physics

H - 1525 Budapest, P.O.Box 49, Hungary

Phone:	+36-1-392-2222/39-62

Fax : 	+36-1-395-9151

Email:	szilagyi@rmki.kfki.hu

G. Amsel, G. Battistig and A. L’Hoir: Nucl. Instr. and Meth. B 201 (2003) 325-388

V. Quillet, F. Abel and M. Schott, Nucl. Instr. and Meth. B 83 (1993) 47.

E. Szilágyi, F. Pászti, A. Manuaba, C. Hajdu, E. Kótai, Nucl. Instr. Meth. B 43 (1989) 502 - 506.

J.F. Ziegler, Nucl. Instr. and Meth. B219–220 (2004) 1027.